## Slide 1
WindCube FPI
Calibration Analysis
Tolansky Two-Line Analysis | n-Pairs Sensitivity | Systematic Error Budget | Future Work Roadmap
Scott Sewell · High Altitude Observatory / NCAR · Boulder, CO
20.107 mm
Etalon gap d
±0.2 µm (1σ)
1.6081×10⁻⁴
Plate scale α
rad/px ±101 ppm
12 → 20+
Ring pairs used
extending to r=150 px
1.77 / 2.86
χ² (line a / b)
improvement path clear

## Slide 2
Error Propagation: What Controls α and d?
Full analytical propagation through the Tolansky chain: r²_fit → WLS slope Δ → Benoit d → α
r²_fit (each ring)
WLS slope Δ (px²/fringe)
Fractional order ε (intercept)
Benoit etalon gap d
Plate scale α
Quantity | Variance share | σ (current)
σ_Δ_a → σ_α | 99.8% ← dominant | 0.25 px²/fr
σ_d → σ_α | 0.2% negligible | ±0.19 µm
σ_ε_b → σ_d | 62% (weak line b) | ±0.0014 fr
σ_ε_a → σ_d | 38% (strong line a) | ±0.0011 fr
✓ d is already excellent
σ_d = ±0.19 µm (9 ppm). Improving ε further has essentially zero impact on σ_α. Don't chase d precision.
⚠ Weak line b is the bottleneck
σ_Δ_b = 0.328 vs σ_Δ_a = 0.25 px²/fr. Line b (638.3 nm) is photon-noise limited throughout. More lamp flux or exposure time helps disproportionately.
→ Path to sub-50 ppm σ_α
Reduce σ_Δ by 50% → σ_α halves. Achievable with brighter neon, more rings, or improved fringe centroiding (Airy fit vs parabola).

## Slide 3
n-Pairs Sensitivity Sweep — Key Findings
Each point re-runs full Tolansky WLS using only the innermost n ring pairs
d is stable across all n
Etalon gap d holds within ±1σ from n=8 to n=12 with no monotonic drift. Outer rings are noisy but unbiased. No systematic error introduced by including them.
σ_d monotonically decreases
d uncertainty falls smoothly from ~0.31 µm at n=4 to ~0.19 µm at n=12, with no sign of flattening. More rings always win for d.
σ_α monotonically decreases
Plate scale uncertainty falls from ~400 ppm at n=4 to ~101 ppm at n=12. Crucially, no inflection point — the outer rings are not biasing α upward despite lower SNR.
Δ slopes are stable
Both Δ_a and Δ_b remain within error bars across all n. No individual ring produces a discontinuous jump, confirming no outlier ring is corrupting the slope.
σ_Δ decreases monotonically
Formal WLS uncertainty on both slopes decreases with every added ring pair. Line b consistently noisier than line a — the photon-noise gap is clear.
χ²_red: the one caveat
χ²_red oscillates 1.5–2.9 for both lines and never settles to 1. Formal σ_Δ values are 35–70% optimistic. True σ_α ~ 130–170 ppm. Systematic error source to investigate.
CONCLUSION: More rings universally improves precision. Outer fringes are noisy but unbiased. Proceed with R_MAX = 150 px (20+ ring pairs).

## Slide 4
Extending Fringe Coverage: ROI → Full Image
Image geometry
Full frame: 259 × 276 px Optical centre (full frame): col=144.7, row=143.7 Min edge distance: 115.3 px (bottom) — limits R_MAX Corner distance: 174–204 px Current ROI R_MAX: 122 px → Extended: 150 px
Peak | r (px) | % ROI | % Full image | Gain
P24/P25 | 122–125 | 37–41% | 87% | +138%
P26/P27 | 128–130 | 28–32% | 85% | +197%
P28/P29 | 133–135 | 22–25% | 76% | +247%
P30/P31 | 138–140 | 16–19% | 70% | +337%
P32/P33 | 143–145 | 11–13% | 67% | +492%
P34+ | 146–155 | < 10% | 40–50% | +600%+
Approach 1 — Already done: R_MAX = 150 px
Annular reduction handles partial annuli correctly. Pixels at each r² are averaged regardless of arc completeness. SEM scales as 1/√n_pixels — larger for partial arcs — so WLS automatically down-weights them. P24–P31 added to _R2_WINDOWS.
Approach 2 — Use the full 259×276 frame
Removes the ROI altogether. Gains 138–1000% more arc coverage for P24+. Boundary artifacts (if confined to a small angular sector) only increase SEM, not bias, as long as they don't systematically offset the mean intensity around the annulus. Test by comparing d and α with and without ROI mask.
Caveat: optical centre offset
Centre is 115 px from the bottom edge vs 145 px from the top. Full-image annuli are asymmetrically clipped. This is fine for photometry but means outer bins have more pixels on one side. The azimuthal mean is still valid; only its uncertainty estimate changes.

## Slide 5
Elevated χ²_red — Candidate Explanations
χ²_red = 1.77 (line a) and 2.86 (line b) — ring positions scatter 35–70% more than formal ±σ_r² predicts
MOST LIKELY
A. Parabola is the wrong local model
The Airy function in r²-space is NOT symmetric — it has a sharper outer flank than inner. A parabola fitted to the top 50% of the peak returns a centroid biased toward the sharper flank. The bias is not constant: it depends on finesse, fringe order, and window placement. With R_refl=0.53 (F≈4.8) fringes are broad and nearly sinusoidal, limiting but not eliminating the bias. Fix: replace parabolic centroid with a full Airy or Gaussian fit.
CONTRIBUTING
B. σ_r² underestimated (formal error too small)
Parabolic centroid covariance assumes Gaussian residuals. Any structured noise — flat-field non-uniformity, etalon plate figure, lab thermal gradients — makes formal σ_r² too small, causing WLS to over-weight those rings. Result: χ²_red > 1 throughout. Fix: rescale σ_r² by √χ²_red for quoted uncertainties (current σ_α ~ 130–170 ppm true vs 101 ppm formal).
IMPORTANT
C. Etalon surface figure errors
Non-flat etalon plates → gap d varies slowly with r. Fringe positions in r² deviate from a perfectly linear Tolansky model. The residual panel B shows a broad bowl shape (negative mid-range, positive at extremes for line a) — classic signature of a slight spherical figure error in one plate. This is an accuracy (not just precision) issue. Fix: characterise plate figure; apply r-dependent d correction.
SECONDARY
D. Thermal / mechanical drift during 30s exposure
Etalon temps logged: [23.63, 23.40, 23.45, 23.46]°C — ~0.2°C spread across sensors. Gap drift during exposure smears fringes and biases inner/outer ring centroids differently. Same effect from vibration. Fix: shorter exposures, active temperature control, or drift-correction using pairs of exposures.
SECONDARY
E. Centre-finding residual error
NM uncertainty σ_cx = σ_cy ≈ 0.02 px. A centring offset δ shifts every r² by ~2r·δ, growing linearly with r. At r=120 px, induced r² error ≈ 5 px² — comparable to ±σ_r² of outer rings. Fix: extend NM optimisation to use a wider search grid; verify centre is consistent across multiple exposures.
CHECK FIRST
F. Incorrect peak family assignment
Peaks are split into line-a vs line-b families by amplitude threshold. If even one peak is mis-assigned, its r² lands at the wrong fringe index p, introducing a step offset in the Tolansky slope. Fix: verify each family by checking predicted vs actual r² spacing; consider assigning by proximity to predicted positions rather than amplitude alone.

## Slide 6
Future Work Roadmap
Prioritised actions to improve α and d precision/accuracy for WindCube science operations
P1
Medium
Replace parabolic centroid with Airy/Gaussian fit
Highest expected χ²_red reduction. Implement as an option in annular_reduce(). Compare centroid positions: systematic offset between parabola and Airy centroid is a direct measure of the current bias.
P1
Low
Brighter neon lamp or longer cal exposure
Line b (638.3 nm, weak) is the dominant σ_Δ bottleneck. Doubling lamp flux or exposure halves photon noise on all rings simultaneously. Most cost-effective single hardware change.
P2
Low
Extend to R_MAX = 150 px (done) then full-frame
P24–P31 already added. Next: remove ROI and use full 259×276 image. Validate by comparing d and α with ROI vs full-frame; if consistent, adopt full-frame as default.
P2
Trivial
Rescale formal uncertainties by √χ²_red
Until χ²_red is understood, multiply all σ_Δ, σ_d, σ_α by √χ²_red before quoting. Currently σ_α ≈ 101 ppm formal but ~130–170 ppm true. This is a bookkeeping fix, not a science improvement.
P3
Medium
Characterise etalon plate figure (candidate C)
Take calibration frames at multiple temperatures. If d shifts with T in a radially-symmetric pattern, it suggests plate figure error. Could also analyse residuals vs r to extract the figure term directly from existing data.
P3
Medium
Verify peak family assignment robustness
Replace amplitude-threshold split with position-prediction assignment: assign each detected peak to the family whose predicted r²_fit (from the running Tolansky slope) is closest. More robust at low SNR.

## Slide 7
WindCube SOC Pipeline — Current Status
20.107 mm
d (etalon gap)
±0.2 µm Benoit/Tolansky
1.6081×10⁻⁴
α (plate scale)
rad/px ±101 ppm (formal)
12 → 20+
Ring pairs
extending outer fringes
1.77 / 2.86
χ²_red (line a/b)
target: < 1.5 — WIP
Pipeline output figures (run_windcube_calibration.py)
Fig 0
Raw images + metadata + dark-sub ROI + histogram
Fig 1
Two-pass azimuthal variance centre finder (NM)
Fig 2
Peak identification table (24 peaks, unified columns)
Fig 3
All-fringe fit diagnostics — 24 sub-panels
Fig 4
Tolansky two-line analysis (Benoit d + plate scale α)
Fig 5
n-pairs sensitivity sweep — precision vs ring count
Scott Sewell · sewell@ucar.edu · HAO/NCAR · soc_sewell repository