## Slide 1
H05 → H06
Calibration Handoff
What the H05 neon inversion delivers to H06 wind retrieval
WindCube FPI · CAL01 · 2026-05-27

## Slide 2
Pipeline context
H05 characterises the instrument. H06 measures the wind.
H05 — Neon calibration
Input: neon lamp fringe image Fits 11-parameter Airy model to r² ≤ 12 000 px² (r ≤ 109.5 px)
Output:
instrument parameter dict
saved to _cal_result.npy
.npy
H06 — Wind retrieval
Input: airglow fringe image + cal_result.npy All instrument params fixed One free parameter: v_LOS
Output:
line-of-sight wind
velocity (m/s)
epsilon_cal anchors the zero-wind phase — every wind measurement is a shift relative to this value

## Slide 3
What .npy contains — field by field
Field
Value (example)
Used by H06?
Notes
t_m
20.1076255 mm
YES — critical
Etalon gap; sets fringe spacing
alpha
1.607e-4 rad/px
YES — critical
Plate scale; maps r → angle
epsilon_cal
0.266318
YES — critical
Zero-wind phase ref; (2t/λ₁) mod 1
R_refl
0.2413 (= R1)
YES
Fringe contrast at OI 630 nm
sigma0
0.5592 px
YES
PSF blur width
I0/I1/I2/I3
1817 / –0.208 / …
YES
Intensity envelope coefficients
B
658 ADU
YES
CCD bias pedestal
R2
0.2830
No
638.3 nm reflectivity; neon only
ne_ratio
0.7883
No
638/640 intensity ratio; neon only
chi2_reduced
1.587
No
Fit quality diagnostic
r2_max_fit
12000 px²
No
Provenance: radial cutoff used
n_pairs_tolansky
16
No
Provenance: Tolansky seed pairs

## Slide 4
The critical link: epsilon_cal
ε_cal = (2 · t / λ₁) mod 1 where λ₁ = 640.2248 nm
What it represents
The fractional interference order at the detector centre when the thermosphere has zero line-of-sight velocity. It is the absolute phase reference.
Why it matters
Without ε_cal, H06 can only measure relative Doppler shifts. With it, shifts are converted to absolute velocities in m/s referenced to the rest wavelength.
How it is used
H06 computes λ_c = λ_OI · (1 + v/c) and finds v by minimising χ² between the observed airglow profile and the model A(r; t, α, R, σ₀, λ_c).
Positive v_LOS (recession) → λ_c > λ₀ → fringes shift inward to smaller r | Negative (approach) → fringes shift outward

## Slide 5
What H06 does not use — and why
R2 (638.3 nm reflectivity)
R2 characterises the neon weak line. The OI 630 nm airglow is a single-wavelength source — only R1 (≈ R_refl at 630 nm) enters the H06 Airy model.
ne_ratio (638/640 intensity ratio)
This is the amplitude ratio of the two neon lamp lines. It has no physical meaning for the airglow source, which emits only at 630.0304 nm.
chi2_reduced, converged, quality_flags
Fit quality diagnostics. H06 reads them only to gate on calibration quality — if chi2_reduced is too high or converged=False, H06 can flag science frames.
r2_max_fit, n_pairs_tolansky, source_file
Provenance fields for traceability and reproducibility. Not consumed in the wind retrieval computation.
t_eff_phase_corrected_mm
The phase-corrected gap used as the LM seed. The fitted t_m is what H06 uses — the seed value is provenance only.